## Slide 1
DEVELOPMENT OF AN EXPERIMENTAL SETUP FOR SAFETY ASSESSMENTS OF NEW ULTRASOUND IMAGING MODES
Adeline Bernard1,
Hervé Liebgott1,
Pauline Muleki-Seya1,
Lorena Petrusca1,
Emmanuel Roux1,
Matthieu Toulemonde2,
François Varray1,
Didier Vray1,
1 Univ Lyon, CREATIS, France
2 Department of Bioengineering, Imperial College London, United Kingdom
This material is based upon work done on the PILoT facility
(PILoT, INSA LYON, Bât. Blaise Pascal, 7 Av. Jean Capelle 69621 Villeurbanne)
Notes: La première diapo fait très chargée ? Puis-je enlever ou diminuer certains logos ?
Penser à ajouter se photo si pas de webcam

## Slide 2: INTRODUCTION - BACKGROUND
Ultrasound medical imaging: Affordable – Real-time / fast imaging – spatial resolution – safety
Open platforms available to research [1] (Verasonics, Ula-Op, …)
Development of new imaging techniques [2] (unfocused, ultrafast, 3D, …)
In vivo acquisition ?
[1] E. Boni, et al. ‘Ultrasound Open Platforms for Next-Generation Imaging Technique Development’, IEEE Trans. Ultrason. Ferroelectr. Freq. Control, 2018,
[2] M. Tanter and M. Fink, ‘Ultrafast imaging in biomedical ultrasound’, IEEE Trans. Ultrason. Ferroelectr. Freq. Control, 2014
Prove that the whole acquisition platform (ultrasound scanner, probes and imaging sequences) is safe and is compliant with safety regulations
MI (Mechanical Index)
ISPTA (Intensity Spatial Peak Temporal Average) and TI (Thermal Index)
Notes: Idea ?

## Slide 3: INTRODUCTION – COMMERCIAL SOLUTIONS
Company | ACERTARA, USA | ONDA CORPORATION, USA | PRECISION ACOUSTIC, UK
Model | AMT | AIM III | UMS3
Image |  |  | 
Cost in k€ (date of quotation) | 48 (2020) | 57 (2020) | 46 (2018)
Automatic calculations | Gold standard | ISPTA, pulse duration, MI (derated value) | ISPTA, pulse duration, MI + others
Regulatory requirements | NEMA UD2 and others | FDA 510(k) | FDA 510(k), NEMA UD2 and others

## Slide 4: INTRODUCTION – COMMERCIAL SOLUTIONS
Company | ACERTARA, USA | ONDA CORPORATION, USA | PRECISION ACOUSTIC, UK
Model | AMT | AIM III | UMS3
Minimum step size in µm | Between 0.198 and 1.25 | 5.5 | 1
Dimension of travel : X x Y x Z in cm | Not found* | 38 x 26,5 x 33 | 30 x 30 x 30
Options | Integrated Thermocouple | Temperature Probe, rotational axes, water conditionning system | rotational axes, water conditionning system
Programming language | Access to the source code possible | External Control via Ethernet or Librairies | Labview + source code for developments
These commercial solutions may not well adapted for new specific ultrasound imaging modes developed in research, for example, for 2D and 3D unfocused imaging.
* This information do not be found on website and datasheet of the product.
Notes: Expliquer le not found ?

## Slide 5: INTRODUCTION – PROBLEM
Focused Imaging
Diagram / simulation [3]
Pressure field obtained with our setup [4]
Plane wave Imaging
[3] S. Shahriari and al, ‘Meshfree simulations of ultrasound vector flow imaging using smoothed particle hydrodynamics’, Phys. Med. Biol., Oct. 2018
[4] M. Toulemonde, ‘Formation de voies en émission et en réception pour l’amélioration de l’imagerie ultrasonore : application à l’imagerie non linéaire des tissus biologiques’, thesis, 2014
Maximum pressure at the depth of focalization
Medium insonify more uniformly
Notes: Est-ce un problème si les références n’apparaissent pas dans le même ordre que dans le proceeding ?

## Slide 6: INTRODUCTION – PROPOSED SOLUTION
Commercial platforms: available with customization possibilities
Proposal: our own pressure field measurement platform
  Perfect control over all possible measurements and calculations.
  Continuously improve our setup according to new standards / requirements.

## Slide 7
Z
Y
X
MATERIALS AND METHOD - DESCRIPTION OF THE EXPERIMENTAL SETUP
A solid bench we already have
An optical breadboard of 1000 x 400 x 33 mm, weight : 32 kg
Three 750 mm profiles mounted in an inverted U-shape and mounting accessories
DC motor-driven translation (Owis, Germany) :
Length X and Z : 95 mm, Y : 195 mm
Accuracy : 1,25 µm
Controller 3 axis with software and DLL
About 10 k€ in 2013

## Slide 8: MATERIALS AND METHOD – PICTURE OF OUR SETUP
SETUP
Water tank with probe and hydrophone
Oscillopscope
Software
Open ultrasound scanner Vantage 256 (Verasonics, USA)

## Slide 9: MATERIALS AND METHOD - COMPARISON OF OUR EXPERIMENTAL SETUP WITH COMMERCIAL SOLUTIONS
 | Our setup | Synthesis of commercial solutions
Architecture | Composed of several parts | One block
Price without options (k€) | 10 (in 2013) | 50 (in average, quotation in 2018-2020)
Automatic calculation of safety indexes | ISPTA, MI and TI | ISPTA, pulse duration, MI + others
Regulatory requirements | IEC60601-2-37 [5], IEC61102 [6] | IEC60601-2-37 + others
Minimum step size in µm | 1.25 | Between 0.198 and 5.5
Dimension of the travel: X x Y x Z in cm | 9.5 x 19.5 x 9.5 | 34 x 28 x 32
Programming language | Labview and Matlab | Customization possibilities
Options | Possible to add | Temperature Probe, rotational axes, water conditionning system
[5] ‘IEC 60601-2-37:2004 Medical electrical equipment – Part 2-37: Particular requirements for the safety of ultrasonic medical diagnostic and monitoring equipment’. International Electrotechnical Commission Edition 1:2001,
[6] I. E. Commission, ‘Measurement and characterization of ultrasonic fields using hydrophones in the frequency range 0.5 MHz to 15 MHz’. IEC, 1991

## Slide 10: MATERIALS AND METHOD – PROCEDURE ACCORDING TO STANDARDS
For non scanning application
Step 1: beam axis
Acquire three 1D scans in the lateral dimension at three depths to find the axis of the ultrasound beam.
Step 2: ISPTA and depth Z3
Calculate ISPTA along the beam axis.
Find the depth Z3 corresponds to the maximum ISPTA.
Step 3: MI
Calculate MI at depth Z3.
Step 4: Beam exit area
Acquire 1D scan in lateral and elevation.
Calculate an area and an associated depth Zbp.
Step 5: Measurement and calculation of the sound power level P3
Acquire a 2D scan at depth Zbp.
Integrate intensities to obtain power P.
Calculate power P3 as a function of the depth.
Step 6 : TI calculation
Calculate TI as a function of P3.

## Slide 11: MATERIALS AND METHOD – DATA ACQUISITION
Acquisition of the complete pressure field
Dimension | Length (mm) | Step (mm)
Axial | 45 | 0.5
Lateral | 46 | 0.5
Elevation | 1 | 1
Plane wave sequence with a Vantage 256 (Verasonics, USA) and a linear probe L7-4 :
Frequency | Number of cycle | Voltage | Frame rate
5.2 MHz | 1 | 50 V | 500 Hz
Pressure field of a plane wave sequence in MPa

## Slide 12: RESULTS – SAFETY INDEX
 | Standard (non scanning procedure) | Complete Field
ISPTA in mW/cm² | 44 | 47
MI | 1.1 | 0.8
TIS / TIB | 0.1/ 0.5 | 
Thresholds given by the FDA regulation [7] for the peripheral vessels:
ISPTA < 720 mW/ cm²
MI < 1.9
[7] U. S. Food and D. Administration, ‘Information for manufacturers seeking marketing clearance of diagnostic ultrasound systems and transducers’, Rockv. MD Cent. Devices Radiol. Health US Food Drug Adm., 2008.

## Slide 13: RESULTS – 3D IMAGING
Acquisition of the pressure field at three depths : 10, 20 and 30 mm
Dimension | Lenght (mm) | Step (mm)
Lateral and elevation | from -10.2 to +10.2 | 0.3 (pitch)
One shoot with the central element of a matrix probe 32 x 32 elements (Vermon, France) with a Vantage 256 (Verasonics, USA) :
Frequency (MHz) | Number of cycle
3 | 4
[8] E. Roux, ‘2D sparse array optimization and operating strategy for real-time 3D ultrasound imaging’, thesis, 2016.

## Slide 14: DISCUSSION AND CONCLUSION
Next steps :
Technological aspect: Add a water conditioner.
Methodological aspect: Decrease the acquisition time vs Increase the resolution.
Regulatory aspect: Consider other standards
A working, open and flexible safety measurement platform
Notes: https://obsproject.com/fr/download

## Slide 15
THANK YOU FOR YOUR ATTENTION
ANY QUESTIONS ?
This work was performed within the framework of the LABEX CELYA (ANR-10-LABX-0060) and the LABEX PRIMES (ANR-11-LABX-0063) of Université de Lyon, within the program "Investissements d'Avenir" (ANR-11-IDEX-0007) operated by the French National Research Agency (ANR)

## Slide 16: More results - MI
Method | Standard / en prenant en compte l’erreur de l’hydrophone | Complete field
MI | 0,75* / 0,844 | 0,81
Elevation position | 5th plan / idem | 5th plan
Axial position | 21 mm / idem | 23,5 mm
Lateral position | -1 mm /idem | -16,5 mm
Min pressure (along the beam axis) | -2,28 Mpa at 24 mm depth / -2,55 Mpa at 24 mm depth | -2,5 Mpa
Min pressure at z3 | -2,19 Mpa / -2,46 MPa | Not concerned
*Error on the proceeding

## Slide 17: More results - ISPTA
Method | Standard / en prenant en compte l’erreur de l’hydrophone/ en prenant en compte la température de l’eau et donc la vitesse et la densité | Complete field
ISPTA mW/cm² | 44,4 / 56,0 / 58,46 | 47,3
Elevation position | 5th plan / idem | 5th plan
Axial position | 21 mm (=z3) / idem | 20 mm
Lateral position | -1 mm / idem | -5,5 mm
Pressure corresponding to the max ISPTA | 4,5 Mpa (along the beam axis and at z3) / 5,07 MPa | 5,2 Mpa
Max pressure | Idem like ISPTA / idem | 5,5 Mpa at 21 mm depth and -5,5 mm in lateral

## Slide 18: MATERIALS AND METHOD – OTHERS DEVICES
This experimental setup could be used with any hydrophone. For this setup, a calibrated capsule hydrophone from Onda (USA) is used:
Capsule -> more ruggedness than needle or membrane
Diameter 200 µm :
  Intermediate sensibility : -265 [dB] re 1V/µPa
  High directivity : to measure the maximum of elements contribution
  Precision < pitch of probes elements
band pass from 1 MHz to 20 MHz : a large panel of frequencies could be tested

## Slide 19: MATERIALS AND METHOD – OTHERS DEVICES
Software :
  Labview : to control the hydrophone displacements and the acquisition of the hydrophone signals by an oscilloscope.
  Matlab : to evaluate the safety indexes from the data acquired.
Probe holders :
CAO file
Holder printed

## Slide 20: MATERIALS AND METHOD - ALIGNMENT
Step 1 : visualization of the hydrophone on the ultrasound scanner and first placement in the center of the probe
Step 2 : A 1D scan performed in the lateral and elevation directions to check the alignment and add, if necessary, an additional length of acquisition in either direction.
Step 3 : realization of a C-scan near the probe and at the depth of interest to control the third misalignment angle.

## Slide 21: ALIGNMENT

## Slide 22: RESULTS – UNCERTAINTIES
Sources of error :
  Hydrophone measures uncertainty : ± 1 dB
  Oscilloscope measures uncertainty
  BNC cable error : neglected
  Error positioning of the motor stage : <25µm / 100 mm
  Quality of the water (not filtered and degassed) : risk of cavitation for pressure > 1 MPa
  Integral calculation under MATLAB software by the trapezium method
Commercial systems must certainly present the same sources of error
Added directly during the conversion between voltage and pressure to have the worst case
The measurement should be repeated several times in order to know the overall error percentage

## Slide 23: Mechanical Index (MI)
With :
fawf : arithmetic-mean acoustic-working frequency
With :
Z3 : cf slide 10
α = 3.10-5 dB/(m Hz) : acoustic attenuation coefficient
Notes: arithmetic mean of the most widely separated frequencies f1 and f2, within the range of three
times f1, at which the magnitude of the acoustic pressure spectrum is 3 dB below the peak
magnitude

## Slide 24: ISPTA

## Slide 25: TIS for non scanning application

## Slide 26: Results in focalization
Pressure field of a focalized sequence in MPa
Acquisition of the complete pressure field
Dimension | Length (mm) | Step (mm)
Axial | 45 | 0.5
Lateral | 46 | 0.5
Elevation | 8 | 0.5
Focused sequence at 35 mm with a Vantage 256 (Verasonics, USA) and a linear probe L7-4 :
Frequency | Number of cycle | Voltage | Frame rate
5.2 MHz | 1 | 50 V | 500 Hz

## Slide 27: RESULTS – SAFETY INDEX
 | Standard (non scanning procedure) | Complete Field
ISPTA in mW/cm² | 122 | 171
MI | 1.3 | 1.6
TIS / TIB | 0.1/ 0.9 | 
Thresholds given by the FDA regulation [7] for the peripheral vessels:
ISPTA < 720 mW/ cm²
MI < 1.9
[7] U. S. Food and D. Administration, ‘Information for manufacturers seeking marketing clearance of diagnostic ultrasound systems and transducers’, Rockv. MD Cent. Devices Radiol. Health US Food Drug Adm., 2008.

## Slide 28: More results - MI
Method | Standard | Complete field
MI | 1.3 | 1.6
Elevation position | 5th plan | 4th plan
Axial position | 17 mm | 36,5 mm
Lateral position | -2 mm | -3 mm
Min pressure (along the beam axis) | -4,07 Mpa at 18,5 mm depth | -6,22 Mpa
Min pressure at z3 | -3,56 MPa | Not concerned
*Error on the proceeding

## Slide 29: More results - ISPTA
Method | Standard | Complete field
ISPTA mW/cm² | 122 | 171
Elevation position | 5th plan | 4th plan
Axial position | 17 mm (=z3) | 36,5 mm
Lateral position | -1,5 mm | 2mm
Pressure corresponding to the max ISPTA | 7,2 Mpa (along the beam axis and at z3) | 11,2 Mpa
Max pressure | Idem like ISPTA | 11,3 Mpa at 37,5 mm depth and 2,5 mm in lateral

## Slide 30: Solutions to explored to decrease the number of acquisitions
Solution 1 : Acquire only the necessary data according to the procedure of the standards -> risk : not taking into account the real maximum of the pressure field
Solution 2 : Optimize the size of the field acquired with an iterative process
Solution 3 : Find the maximum pressure using a gradient algorithm.

## Slide 31: Pourquoi le max d’ISPTA ne correspond pas au max de pression?
Max d’ISPTA : pressure(:,31,36,5)
Max de pression : pressure(:,33,36,5)
PPI = 2.8536e+06
PII3= 0.9034
On multiplie PPI par 10^(-alpha*pas_ax(31)*0.001*fawf/10), soit par 0.4875
PPI = 2.9528e+06
PII3= 0.9018
On multiplie PPI par 10^(-alpha*pas_ax(31)*0.001*fawf/10), soit par 0.4703
Cela vient du terme en 10^ car le max de pression est un peu plus profond